## Introduction
PSA OPEX Estimator
An Advocacy and Planning Tool for PSA Plants
INSTRUCTIONS: Read through the Introduction tab for information about the tool.
TOOL PURPOSE
This tool is designed to help global health stakeholders, procurers, and advocates estimate the annual operating expenditures (OPEX) of pressure swing adsorption (PSA) oxygen plants in low- and middle-income countries (LMICs). It is intended to support advocacy, planning, and comparative analysis related to oxygen system investments. The tool generates approximate operating cost estimates based on key PSA plant specifications and operating assumptions. Users should complete all input fields to the best of their ability using information available for the plant being assessed. Suggested values are provided for selected inputs to support users when site-specific information is unavailable. The estimates produced by this tool are intended to provide a standardized, evidence-informed basis for discussion and comparison across settings. The tool is not intended to serve as a detailed budgeting, procurement, or financial planning tool.
TOOL OVERVIEW
Users enter information about the PSA plant being assessed in the OPEX Estimator tab. All fields requiring user input are highlighted in green and include instructions to guide data entry. Gray-shaded fields are also editable; however, these fields are prepopulated with standard values that are generally appropriate for most assessments and are unlikely to require modification. The worksheet is organized into two main sections: Data Entry and Results. The Data Entry section is divided into two parts. The first, Configuration, captures basic information about the PSA plant, including its size, technical specifications, and utilization. The second, Operating Assumptions, captures key operational details such as staffing, maintenance practices, and energy costs. Users should complete the Plant Configuration section first, as many of the suggested values and recommendations in the Operating Assumptions section depend on these inputs. Users are encouraged to review all entered values against the suggested values provided by the tool and to verify all inputs before interpreting the results. The Results section presents estimated annual operating costs by major cost category, including staffing, repair and maintenance, and energy. It also reports the plant's estimated annual oxygen output in both normal cubic meters (Nm³) and oxygen cylinders of the size specified in the Plant Configuration section. The Results section also includes a cylinder procurement comparison. Users may enter the local market cost of procuring a cylinder of equivalent size, and the tool will estimate the annual cost of purchasing the same number of cylinders as the PSA plant is expected to produce over a year. This comparison is intended to illustrate the scale of oxygen production and support advocacy discussions regarding oxygen supply investments.
KEY CONSIDERATIONS
Assess multiplex systems individually. If the facility operates a multiplex PSA configuration with multiple PSA units, estimate each system separately. Operating costs, utilization, and maintenance requirements may differ across units, and assessing them individually will generally produce more accurate results. Consider local operating conditions carefully. Operating costs can vary substantially by country, region, and facility. Users should pay particular attention to assumptions related to labor, maintenance, and energy costs, and adjust default values where reliable local information is available. Interpret estimates cautiously at extreme utilization levels. Estimates for plants operating at very low utilization (e.g., 1–2 hours per day) or very high utilization (e.g., more than 20 hours per day) may not fully capture long-term operational effects. These may include accelerated equipment wear, increased maintenance requirements, shortened component lifespans, or deterioration associated with extended idle periods. Results are sensitive to input assumptions. Small changes in key assumptions, particularly staffing levels, electricity costs, utilization rates, and maintenance expenditures, can have a meaningful impact on estimated annual operating costs. Users should review all inputs carefully before interpreting results. Use local technical knowledge whenever available. Where plant-specific information is known, it should be used in place of suggested values. Suggested values are intended as reasonable defaults when local data are unavailable, not as universally applicable standards.
LIMITATIONS
The cost estimates presented in this tool are intended to capture the major operating and maintenance costs associated with PSA oxygen production and cylinder filling but may not represent all expenditures incurred by a facility. Depending on the site context, additional costs may include: Tools, Supplies, and Personnel Safety: Replacement tools and equipment; cleaning supplies (including mops, paper towels, and spray detergent); and personal protective equipment (PPE), including goggles, ear plugs, electrical gloves, and work clothes. Oxygen System and Fire Safety Maintenance: Maintenance of fire prevention systems, including fire extinguishers; cylinder accessories, including cylinder caps, pigtails, and adapters; maintenance of oxygen distribution pipelines and associated accessories. Facility and Site Maintenance: Facility and plant house maintenance, including air conditioning maintenance; maintenance of security and safety measures, including fences, alarms, and cameras. Cylinder Transportation Operations: Operating costs associated with cylinder transportation and distribution, including fuel, vehicle parts, inspections, and related expenses. Users should review these and other site-specific cost categories to determine whether additional expenditures should be incorporated into the total operating cost estimate.
QUESTIONS
For questions regarding field definitions, data requirements, or any other aspect of the PSA OPEX Estimator, please contact Dr. Alex Rothkopf, Data Science and Supply Chain Advisor, PATH (arothkopf@path.org) or Madison Bradley, Data Science Senior Associate, PATH (mbradley@path.org).

## OPEX Estimator
PSA OPEX Estimator
An Advocacy and Planning Tool for PSA Plants
DATA ENTRY | [7 empty] | Results to the right >> | RESULTS
| PSA Plant Inputs | [7 empty] | Annual Operating Expense (OPEX) Estimate
|  | Configuration |  |  | Operating Assumptions |  |  |  |  | Annual OPEX
|  | 1. Plant Capacity |  |  | 1. Staffing | [5 empty] | Staff | 21300
[11 empty] | Biomedical Engineer(s) | 300
|  | Installed Capacity |  |  | FTE | [5 empty] | Plant Operations Supervisor(s) | 3000
|  | Enter the installed capacity for the PSA plant. If the plant is a multiplex system, estimate each sub-system separately when the units have different functions or operating schedules. If the multiplex system operates as a single integrated plant and the sub-systems are always run simultaneously, the multiplex may be treated as one system, and the installed capacity entered should reflect the total capacity of all units combined. |  |  | Enter the number of full-time equivalent (FTE) staff required for each listed position. One FTE represents one individual working a full-time schedule. Use the suggested values provided in column K as a starting point, which are based on the given Configuration, and adjust as needed to reflect site-specific staffing requirements. | [5 empty] | Plant Operator(s) | 4000
[11 empty] | Cylinder Operator(s) | 14000
[11 empty] | Repair and Maintenance | 58000
[11 empty] | Plant Maintenance | 23000
[11 empty] | Generator Maintenance | 35000
[5 empty] | Biomedical Engineer | 0.05 | Suggested Value: 0.05 FTEs |  |  |  | Solar Maintenance | 0
|  | Installed Capacity | 25 | Nm3/h | Plant Operations Supervisor | 0.6 | Suggested Value: 0.6 FTEs |  |  |  | Energy | 89000
[5 empty] | Plant Operator(s) | 1.73 | Suggested Value: 1.73 FTEs |  |  |  | Primary Source | 47000
|  | Distribution Method |  |  | Cylinder Operator(s) | 6 | Suggested Value: 6.00 FTEs |  |  |  | Secondary Source | 42000
|  | Consider the method by which oxygen leaves the system and is delivered to patients. Enter the proportion of operating time the system directs oxygen to piping and the proportion of operating time used to fill cylinders. In most LMIC contexts, these two proportions should sum to 100%. | [8 empty] | TOTAL | 168300
Unit Costs (monthly, gross)
[5 empty] | Enter the monthly gross salary for each listed position. Suggested values in column K are LMIC market averages; adjust as needed to reflect site-specific conditions. | [5 empty] | Annual Oxygen Output | 219000 | m3
[11 empty] | Or, in Terms of Cylinders of Specified Size | 32205 | cylinders
|  | Proportion of Piping Time | 50 | % | Biomedical Engineer | 500 | Suggested Value: $500 |  | COMPARE
|  | Proportion of Cylinder-Filling Time | 50 | % | Plant Operations Supervisor | 350 | Suggested Value: $350
[5 empty] | Plant/Cylinder Operator | 200 | Suggested Value: $200
|  | Cylinders | [7 empty] | Purchasing Cylinders
|  | Enter the oxygen volume of the cylinders filled by the PSA plant, expressed in Nm³. As a guide, common cylinder sizes are approximately 0.34 Nm³ (Type D), 0.68 Nm³ (Type E), 1.36 Nm³ (Type F), 3.4 Nm³ (Type G), and 6.8 Nm³ (Type J) based on 2020 guidance by the World Health Organization.1 Next, enter the maximum cylinder filling rate, defined as the number of cylinders of this size that can be filled per hour when the PSA plant's full output is dedicated to cylinder filling. |  |  | 2. Repair and Maintenance |  |  |  |  | Enter the cost of a cylinder of the specified size in the local market.
[5 empty] | Plant and Filling Station Maintenance | [5 empty] | Cost Per Cylinder | 0
Enter the annual cost for spare parts and maintenance labor for the PSA plant and associated filling station (if applicable). Maintenance labor should include both preventive and corrective maintenance activities. Suggested values provided in column K are averages based on the installed capacity specified under Configuration. Costs may vary significantly by plant size, manufacturer, and operating conditions. Users are encouraged to validate these estimates with suppliers or maintenance providers whenever possible. Suppliers often suggest budgeting approximately 10% of the PSA plant capital cost annually for spare parts, preventive maintenance, and corrective maintenance combined.
[11 empty] | Annual Cost to Procure | 0
|  | Cylinder Size | 6.8 | Nm3 |  |  |  |  | REFERENCES
[9 empty] | 1 | World Health Organization. Priority Medical Devices List for the COVID-19 Response and Associated Technical psecifications. World Health Organization; 2020. https://www.who.int/publications/i/item/WHO-2019-nCoV-MedDev-TS-O2T.V2
|  | Max Cylinders Filled per Hour | 3.68 | Suggested Value: 3.68 cylinders
|  | 2. Technical Specifications | [6 empty] | 2 | International Finance Corporation. The Dirty Footprint of the Broken Grid. International Finance Corporation; 2019. https://www.ifc.org/content/dam/ifc/doc/mgrt/20190919-full-report-the-dirty-footprint-of-the-broken-grid.pdf
|  | Energy Requirements |  |  | Annual Spare Parts | 12500 | Suggested Value: $12,500
|  | Enter the total rated power for the PSA plant. This should reflect the sum of all core system components (e.g., air compressor, dryer, oxygen generator, and other integrated equipment), though a booster (if present) should be listed separately. |  |  | Annual Maintenance Labor | 10000 | Suggested Value: $10,000 |  | 3 | Feldman D, Ramasamy V, Fu R, Ramdas A, Desai J, Margolis R. U.S. Solar Photovoltaic System and Energy Storage Cost Benchmark: Q1 2020. National Renewable Energy Laboratory; 2021. https://docs.nlr.gov/docs/fy21osti/77324.pdf
Generator Maintenance
[5 empty] | Enter the generator maintenance cost coefficient in dollars per kWh ($/kWh). This value represents the maintenance cost per unit of electricity generated and will be applied proportionally based on system operation. A value of $0.40/kWh is recommended based on a 2019 study conducted by the International Finance Corporation.2 |  |  |  | 4 | Xie H, Ringler C, Mondal A. Solar or diesel: a comparion of costs for groundwater-fed irrigation in Sub-Saharan Africa under two energy solutions. Earth's Future. 2021;9(4). https://doi.org/10.1029/2020EF001611
PSA Plant
|  | PSA Plant Power | 30 | kW
[9 empty] | 5 | Center for Sustainable Systems, University of Michigan. Photovoltaic Enegy Factsheet. Center for Sustainable Systems, University of Michigan; 2026. https://css.umich.edu/sites/default/files/2026-08/CSS07-08_0.pdf
Booster
|  | Booster Power (0 if none) | 7 | kW | Generator Maintenance Coefficient | 0.4 | $/kWh
|  | 3. Operating Conditions |  |  | PV Maintenance
Enter the PV maintenance coefficient in dollars per kWp per year ($/kWp/year). The value represents the maintenance cost per unit of installed solar capacity and will be scaled based on the system size specified under Configuration. A value of $20/kWp/year is recommended based on a 2021 study by the National Renewable Energy Laboratory.3
Operating Hours
Enter the average number of hours the PSA plant operates per day. Use a representative value that reflects typical operating conditions over time (e.g., accounting for downtime, maintenance, or variability in demand) rather than a maximum or theoretical value.
[5 empty] | PV Maintenance Coefficient | 20 | $/kWp/year
|  | Average Daily Operating Hours | 24 | hours
3. Energy Costs
Annual Operation
|  | Enter the number of days per year that the PSA plant is expected to operate. Use a realistic estimate that reflects planned shutdowns, maintenance periods, and any unexpected downtime. |  |  | Energy Prices
Enter the cost of grid electricity and diesel in dollars per unit. Country-level average prices can be referenced at www.globalpetrolprices.com; adjust as needed to reflect site-specific rates or contractual agreements.
|  | Days of Operation | 365 | days per year
[5 empty] | Grid Electricity Cost | 0.23 | $/kWh
|  | 4. Supporting Infrastructure |  |  | Diesel Cost | 1.18 | $/L
|  | Solar |  |  | Energy Efficiency
|  | Consider whether the PSA plant is supported by solar power. If yes, enter the installed capacity of the solar PV system in kilowatt-peak (kWp). For sites with a large shared solar installation, estimate and enter only the portion of total capacity allocated to the PSA plant. The suggested value assumes approximately four peak sun hours per day (i.e., periods of maximum solar generation). Adjust the entered capacity as needed if local solar conditions differ materially from this assumption. |  |  | Enter the generator efficiency in liters per kWh (L/kWh) and the PV system efficiency as a percentage (%). These values represent the fuel consumption per unit of electricity generated (generator) and the conversion efficiency of solar panels (PV), respectively. A value of 0.40 L/kWh is recommended for generators based on a 2021 study by Xie, Ringler, and Mondal.4 A value of 20% is recommended for PV systems based on a 2026 University of Michigan solar factsheet.5 Adjust as needed to reflect site-specific equipment performance and operating conditions.
|  | Solar Rating (0 if none) | 0 | kWp
|  |  | Suggested Value: 0 kWp |  | Generator Efficiency | 0.4 | L/kWh
[5 empty] | PV Efficiency | 20 | %
Energy Mix
Specify the primary and secondary energy sources used to operate the PSA plant. Enter the corresponding code for each source (1 = grid electricity, 2 = diesel, 3 = solar). Provide the proportion of total energy supplied by the primary source (expressed as a share of the total—e.g., a percentage). The proportion for the secondary source will be automatically calculated as the remainder.
[5 empty] | Primary Energy Source | 1
[5 empty] | % of Energy Needs Met by Primary | 70%
[5 empty] | Secondary Energy Source | 2
[5 empty] | % of Energy Needs Met by Secondary | 30%